Significant risk mitigation action plan

Name of healthcare service:

Facility ID:

Accrediting agency:

Address:

Main contact person:

Purple text is for guidance only, delete before use. Refer to Advisory PCHS26/03: Notification of significant risk for full information.

| Healthcare service to complete within 48 hours of risk identification |  |  |  |  |  |  | Healthcare service to track progress and complete before final assessment | Accrediting agency to complete at final assessment |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Item # | Description of the significant risk | Date of identification | Initial risk rating | Actions required | Person responsible | Follow up risk rating | Progress notes | Date of final assessment | Evidence | Final risk rating |
| 1 | - The infection prevention and control manual is in its template state and has not been tailored to the practice - Staff demonstration of reprocessing processes was not consistent with relevant national and international standards - Steriliser calibration has not been conducted in the last three years - No evidence of daily steriliser maintenance - No evidence of patient-level tracking of sterilised instruments. | 01/10/2025 | Extreme | - Immediately stop reprocessing instruments on site and only use single-use instruments - Review and complete the infection prevention and control manual by 15/10/2026 - Appoint a person responsible for infection prevention and control to ensure oversight and compliance by 15/10/2026 - Conduct staff training on the reprocessing process and equipment operation and maintenance as soon as possible - Conduct steriliser validation and calibration as soon as possible then at least annually and whenever otherwise required by the manufacturer or standards - Conduct daily steriliser maintenance - Maintain ongoing documentation and records. | Principal dentist and practice manager | Medium | 01/10/2026-15/01/2026 - Stopped reprocessing instruments on-site and purchased single-use instruments 01/10/2026 - Completed the infection prevention and control manual 15/10/2026 - Completed all staff meeting to discuss infection and control procedures, and appointed the person responsible on 15/10/2025 - Completed staff training on 03/11/2025 - Completed steriliser validation and calibration on 03/11/2025, set up annual reminder - Resumed the use of steriliser on 03/11/2025 - Ongoing maintenance of daily sterilisation maintenance log - Ongoing patient-level tracking of sterilised instruments. | 16/01/2026 | - Reviewed the updated infection control prevention and control manual - Reviewed relevant meeting minutes and other documents relating to steriliser maintenance and staff training - Inspected the sterilisation station - Verified staff understanding of the updated infection prevention and control manual, policies and procedures through interviews and a walk-through of the process. | Low |